## Slide 1: Low-cost, open-source, flexible current feedback-controlled coil driver circuit for matrix B0 shim coils and related applications
Nicolas Arango1, Jason P Stockmann2, Thomas Witzel2,3, Lawrence L Wald2,3, Jacob White1
1 Department of Electrical Engineering & Computer Science, Massachusetts Institute of Technology, Cambridge, MA, USA
2 Martinos Center for Biomedical Imaging, Massachusetts General Hospital, Charlestown, MA, USA
3 Harvard Medical School, Boston, MA, USA

## Slide 2
Nick Arango

## Slide 3: Original Motivation: Flexible local-multicoil B0 shim array driver
Agilent E3648A DC supply
Yale modular driver array
MGH First Gen. driver
(courtesy Thomas Witzel)
Cost of one-off shim drivers
Non-recoverable engineering time
Impedes reproducibility
Too costly for classroom use
Distracts researchers from MR physics and clinical applications
Juchem et al., JMR 2010
Truong et al., NeuroImage 2014
Stockmann et al., MRM 2015
Diversity of multi-coil shim arrays:

## Slide 4: Low-cost and available hardware
Low cost, commodity parts, under $70 per channel
One day assembly by technician or turnkey assembly from multiple vendors
Rapid rise time for dynamic shimming of single and multi-turn local shim coils
Low distortion driving resonating coils
Stable low RF noise for use in the magnet room
Notes: Low cost
Capable of shimming Waveforms driving single and multiturn local shim coils and resonating coils
Low noise and usable in magnet room

## Slide 5: Methods: Linear, current feedback, lead-compensated push-pull
Gain (dB)
Phase (deg)
Frequency (kHz)
Lead Compensated Open Loop Bode Plot
49.5 degree phase margin
~20KHz bandwidth
Single channel

## Slide 6: Results: High performance in diverse applications
Driving 10uH load
20𝜇s rise time
Stockmann et al., MRM 2015
Juchem et al., JMR 2010
Driving 690uH load
200𝜇s rise time
Driving 25KHz resonant load
50A peak to peak
Cooley et al., ISMRM 2014
50uH load
Measured response
Calculated response
Input step

## Slide 7: Results: High performance in diverse applications
Driving 10uH load
20𝜇s rise time
Stockmann et al., MRM 2015
Juchem et al., JMR 2010
Driving 690uH load
200𝜇s rise time
Driving 25KHz resonant load
50A peak to peak
Cooley et al., ISMRM 2014
50uH load
Measured response
Calculated response
Input step

## Slide 8: Results: High performance in diverse applications
Driving 10uH load
20𝜇s rise time
Stockmann et al., MRM 2015
Juchem et al., JMR 2010
Driving 25KHz resonant load
50A peak to peak
Cooley et al., ISMRM 2014
50uH load
Measured response
Calculated response
Input step
Driving 690uH load
200𝜇s rise time

## Slide 9: Results: Simple changes allow for new applications
load
Two capacitor per channel adjustment for new loads
Gain (dB)
Phase (deg)
Frequency (kHz)
Lead Compensated Open Loop Bode Plot
~20KHz bandwidth

## Slide 10: Results: Simple changes allow for new applications
Gain (dB)
Phase (deg)
Frequency (kHz)
Lead Compensated Open Loop Bode Plot
~20KHz bandwidth
load
Two capacitor per channel adjustment for new loads
load
Parallel outputs for larger loads

## Slide 11: Results: Dynamic disturbance rejection
Sequence parameters: EPI GRE sagittal, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
3T scanner bore
15 cm
from isocenter
Position test loop 15 cm from isocenter and oriented for maximum back-EMF caused by EPI readout (z) gradient slewing to maximize flux
Notes: READOUT ALONG Z SAY THAT THING creating a large amo

## Slide 12: Results: Dynamic disturbance rejection
Sequence parameters: EPI GRE sagittal, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
3T scanner bore
Position test loop 15 cm from isocenter and oriented for maximum back-EMF caused by EPI readout (z) gradient slewing to maximize flux
Gradient slew disturbance

## Slide 13: Results: Dynamic disturbance rejection
Sequence parameters: EPI GRE sagittal, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
3T scanner bore
Shim current during gradient switching
Closed loop controller rejects gradient slew

## Slide 14: Results: Dynamic disturbance rejection
Sequence parameters: EPI GRE sagittal, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
3T scanner bore
Shim current during gradient switching
Closed loop controller rejects gradient slew

## Slide 15: Results: Shim current can be set 500 𝜇s before image TR
Sequence parameters: EPI GRE axial, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
+2A
0A
-2A
0A
Single-turn test loop
Phantom
3T scanner bore
Signal ROI
Noise ROI
Test: Update shims to +/-2A starting 500us before beginning of TR in EPI-GRE sequence
Current driver
EPI magnitude images (20 slices)
a.u.
shim current OFF
SNR = 48.0
shim current ON
SNR = 48.5
Notes: Shows correct distortion in the phase encode direction

## Slide 16: Results: No excess RF noise caused by dynamic shimming
Sequence parameters: EPI GRE axial, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
Phantom
3T scanner bore
Test: Update shims to +/-2A starting 500us before beginning of TR in EPI-GRE sequence
→ No RF “zipper” artifacts or structured noise
a.u.
+2A
0A
-2A
0A
EPI magnitude images (20 slices)
shim current OFF
SNR = 48.0
shim current ON
SNR = 48.5

## Slide 17: Future Work: Driver board as a part of a shimming system
Low Cost Microcontroller
Fiber Optic Driver
SPI Bus
Fiber Optic Receiver
8 ch. shim driver #1
RF Waveguide
Magnet room
Equipment/console room
Scanner Console
TTL Pulse
SPI Bus
8 ch. shim driver #2
8 ch. shim driver #7
8 ch. shim driver #8
64 channel update in under 500 𝜇s
64 channel update in under 500 microseconds

## Slide 18: Conclusions
Showed open source, low-cost current driver board tailored for multi-coil shimming (up to 5A DC)
Under $70 per channel
Demonstrated dynamic shimming ability
Stable currents under heavy gradient slewing (EPI)
Low noise operation in magnet room
Limitation: Requires water cooling for high duty cycles
Boards and simulation codes available at rflab.martinos.org

## Slide 19: Conclusions
Showed open source, low-cost current driver board tailored for multi-coil shimming (up to 5A DC)
Under $70 per channel
Demonstrated dynamic shimming ability
Stable currents under heavy gradient slewing (EPI)
Low noise operation in magnet room
Limitation: Requires water cooling for high duty cycles
Boards and simulation codes available at rflab.martinos.org
Notes: “IN conclusion”

## Slide 20: Conclusions
More user-friendly Rev B board forthcoming. Application and feature suggestions welcome
Showed open source, low-cost current driver board tailored for multi-coil shimming (up to 5A DC)
Under $70 per channel
Demonstrated dynamic shimming ability
Stable currents under heavy gradient slewing (EPI)
Low noise operation in magnet room
Limitation: Requires water cooling for high duty cycles
Boards and simulation codes available at rflab.martinos.org
Notes: “IN conclusion”

## Slide 21: Acknowledgements
Grant support: Skoltech
Grant support: NIH R21 EB017338 and P41 EB015896
Notes: I thank my sponsors for funding and you for your attention